	METHOD STATEMENT
	[image: image1.jpg]N







Determinand:
Vinyl chloride (chloroethene)
Matrix:

Sample Type: potable and raw waters
Principle of Method:

The water sample is placed in a septum vial and allowed to equilibrate with its headspace vapour at 60°C.  
A sample of the vapour is injected using an automatic headspace sampler into a capillary column gas chromatograph (GC), the volatile organic compounds are separated and then identified and quantified with mass spectrometric detection (MSD) in selected ion monitoring (SIM) mode.
Sampling and Sample Preparation:

Amber sampling vials are prepared prior to sending to samplers by the addition of 300µl of 1% (W/V) Sodium Thiosulphate solution as specified in the sampling manual.

The vial is slowly and completely filled to exclude headspace in order to avoid loss of volatile determinands. When the water just begins to overflow the vial it is capped with the PTFE face of the septum in contact with the sample. 
Samples are assessed for suitability on arrival in the laboratory, by inversion of the vial, where the bubble covers approximately two thirds of the bottom of the vial this is approximately equivalent to a bubble of 2 ml. Any bubble larger than this (i.e. covers the full width of the vial) is deemed to be significant, and a deviating comment should be made against the relevant test.
Storage – samples should be analysed as soon as possible after collection. When this is not possible, they should be stored under refrigeration at 3 ± 2°C in the dark, until analysis can begin. The maximum permissible storage time prior to analysis is 28 days, which is the maximum period of analyte stability applied to drinking water, raw ground water, and raw surface water matrices
Interferences
Any compound, which passes through the extraction procedure, and has similar Gas Chromatographic and mass spectrometric properties to the analyte will cause interference.
Performance of Method:

Range of Application:

LOQ to 2.500 µg/l
Limit of Quantification:

	Determinand
	LOQ GCHS1 (μg L-1)
	LOQ GCHS2 (μg L-1)
	LOQ Method Standardised 

(μg L-1)

	Vinyl Chloride
	0.113
	0.076
	0.113


Recoveries of Compounds and  Uncertainty of measurement:

Instrument 1: WGCHS1
	Determinand
	Limit of Quantification
(ng L-1)
	Uncertainty of Measurement
(UoM) (%)
	Direct Standards
	Elvington Treated Water 

	
	
	
	Low Standard
	High Standard
	PCV Spike

	
	
	
	Recovery
	RSD
	Recovery
	RSD
	Recovery
	RSD

	Vinyl Chloride
	113
	44.145
	100.80
	5.00
	90.96
	5.96
	99.33
	5.15


Instrument 2: WGCHS2

	Determinand
	Limit of Quantification
(ng L-1)
	Uncertainty of Measurement
(UoM) (%)
	Direct Standards
	Elvington Treated Water 

	
	
	
	Low Standard
	High Standard
	PCV Spike

	
	
	
	Recovery
	RSD
	Recovery
	RSD
	Recovery
	RSD

	Vinyl Chloride
	76
	32.30
	100.01
	2.97
	95.75
	2.01
	91.81
	2.86


References:

Vinyl Chloride in Drinking Water, World Health Organisation, WHO/SDE/WSH/03.04/119
Air Quality Guidelines, Second Edition, Chapter 5.16, Vinyl Chloride, World Health Organisation Regional Office for Europe, Copenhagen, Denmark, 2000.

Vinyl Chloride, Public Health Statement, Agency for Toxic Substances and Disease Registry (ASTSDR), Division of Toxicology and Environmental Medicine, July 2006.

Vinyl Chloride - ToxFAQs, Agency for Toxic Substances and Disease Registry (ASTSDR), Division of Toxicology and Human Health Sciences, July 2006.

NA-TM-1102 v03 VOC/BTEX/F1/VH by HS-GCMS-FID, National Test Method, ALS Environmental (Canada).
	WPC63
	Last Updated: 22/05/26
	Page 2 of 2



[image: image1.jpg]