	METHOD STATEMENT
	[image: image1.jpg]N







Determinand:
The determination of selected of triazine, conazole, phenyl urea, carbamate, organophosphorus and neutral compounds
Matrix:

Sample Type: Treated and Raw Waters
Principle of Method:

A direct aqueous large volume injection (LVI) on-line solid phase extraction procedure. Samples are analysed by high performance liquid chromatography with a triple quadrupole mass spectrometer detector. Samples are injected by large volume injection onto an enrichment column, the enrichment column is back flushed on to the analytical column. Organic compounds are separated and then identified and quantified with mass spectrometric detection in selected reaction monitoring (SRM) mode. Quantitation is based on an internal standardisation procedure.
Sampling and Sample Preparation:

Sampling, samples should be collected in 500mL coloured glass which has been proven to be suitable for this analysis, with PTFE lined screw caps and contain 0.500mL of sample preservative, 30g L-1 w/v ascorbic acid.

Storage – samples should be analysed as soon as possible after collection. When this is not possible, they should be stored under refrigeration at 1-5°C in the dark, until analysis can begin.
The maximum permissible storage time prior to analysis is given below:
	Determinand
	Maximum period of analyte stability prior to any extraction step (days) Drinking water
	Maximum period of analyte stability prior to any extraction step (days)
Surface water
	Maximum period of analyte stability prior to any extraction step (days)
Ground water
	Maximum period of analyte stability after to any extraction step (days)
	Data is quoted from BS EN ISO 5667-3: 2018 [“ISO”] or ALS in-house data [“ALS-AS IHD”]

	ATRAZINE
	21
	21
	21
	N/A
	ALS IHD

	ATRAZINE-DESETHYL
	21
	21
	21
	N/A
	ALS IHD

	ATRAZINE-DESISOPROPYL
	21
	21
	21
	N/A
	ALS IHD

	AZINPHOS-METHYL
	21
	21
	21
	N/A
	ALS IHD

	AZOXYSTROBIN
	22
	21
	21
	N/A
	ALS IHD

	BENZOVINDIFLUPYR
	21
	21
	21
	N/A
	ALS IHD

	BIXAFEN
	22
	21
	21
	N/A
	ALS IHD

	BOSCALID
	21
	21
	21
	N/A
	ALS IHD

	BROMACIL
	21
	21
	21
	N/A
	ALS IHD

	CARBENDAZIM
	21
	21
	21
	N/A
	ALS IHD

	CARBETAMIDE
	21
	21
	21
	N/A
	ALS IHD

	CHLORFENVINPHOS
	21
	21
	21
	N/A
	ALS IHD

	CHLORTOLURON
	21
	21
	21
	N/A
	ALS IHD

	CYANAZINE
	21
	21
	21
	N/A
	ALS IHD

	CYPROCONAZOLE
	21
	21
	21
	N/A
	ALS IHD

	DICHLORVOS
	21
	10
	21
	N/A
	ALS IHD

	DICLOFENAC
	21
	21
	21
	N/A
	ALS IHD

	DIFENOCONAZOLE
	21
	21
	21
	N/A
	ALS IHD

	DIMETHENAMID-P
	21
	21
	21
	N/A
	ALS IHD

	DIMETHOATE
	21
	21
	21
	N/A
	ALS IHD

	DIURON
	21
	21
	21
	N/A
	ALS IHD

	EPTC
	21
	21
	21
	N/A
	ALS IHD

	EPOXICONAZOLE
	21
	21
	21
	N/A
	ALS IHD

	FLUFENACET
	21
	21
	21
	N/A
	ALS IHD

	FLUOPYRAM
	21
	21
	21
	N/A
	ALS IHD

	FLURTAMONE
	21
	21
	21
	N/A
	ALS IHD

	FLUSILAZOLE
	21
	21
	21
	N/A
	ALS IHD

	FLUTRIFOL
	21
	21
	21
	N/A
	ALS IHD

	FLUXAPYROXAD
	21
	21
	21
	N/A
	ALS IHD

	ISOPROTURON
	21
	21
	21
	N/A
	ALS IHD

	LENACIL
	21
	21
	21
	N/A
	ALS IHD

	LINURON
	21
	21
	21
	N/A
	ALS IHD

	MALATHION
	21
	10
	21
	N/A
	ALS IHD

	MEFENTRIFLUCONAZOLE
	21
	21
	21
	N/A
	ALS IHD

	METAMITRON
	21
	21
	21
	N/A
	ALS IHD

	METAZACHLOR
	21
	21
	21
	N/A
	ALS IHD

	METCONAZOLE
	22
	21
	21
	N/A
	ALS IHD

	METHABENZTHIAZURON
	21
	21
	21
	N/A
	ALS IHD

	METHIOCARB
	21
	10
	21
	N/A
	ALS IHD

	METOBROMURON
	21
	21
	21
	N/A
	ALS IHD

	METOXURON
	21
	21
	21
	N/A
	ALS IHD

	METRIBUZIN
	21
	21
	21
	N/A
	ALS IHD

	METSULFURON-METHYL
	14
	21
	21
	N/A
	ALS IHD

	MEVINPHOS
	21
	21
	21
	N/A
	ALS IHD

	MONURON
	21
	21
	21
	N/A
	ALS IHD

	OXADIXYL
	21
	21
	21
	N/A
	ALS IHD

	OXAMYL
	21
	9
	10
	N/A
	ALS IHD

	PROCHLORAZ
	22
	21
	21
	N/A
	ALS IHD

	PROMETRYN
	21
	21
	21
	N/A
	ALS IHD

	PROPACHLOR
	21
	10
	21
	N/A
	ALS IHD

	PROPAZINE
	21
	21
	21
	N/A
	ALS IHD

	PROPETAMPHOS
	21
	21
	21
	N/A
	ALS IHD

	PROPICONAZOLE
	21
	21
	21
	N/A
	ALS IHD

	PROSULFOCARB
	21
	21
	21
	N/A
	ALS IHD

	PROTHIOCONAZOLE-DESTHIO
	22
	21
	21
	N/A
	ALS IHD

	PYRACLOSTROBIN
	21
	21
	21
	N/A
	ALS IHD

	QUINMERAC
	21
	21
	21
	N/A
	ALS IHD

	SIMAZINE
	21
	21
	21
	N/A
	ALS IHD

	SPIROXAMINE
	21
	14
	21
	N/A
	ALS IHD

	TEBUCONAZOLE
	21
	21
	21
	N/A
	ALS IHD

	TERBUTHYLAZINE
	21
	21
	21
	N/A
	ALS IHD

	TERBUTRYN
	21
	21
	21
	N/A
	ALS IHD

	THIFENSULFURON-METHYL
	21
	21
	21
	N/A
	ALS IHD

	TRIAZOPHOS
	21
	21
	21
	N/A
	ALS IHD

	TRIETAZINE
	21
	21
	21
	N/A
	ALS IHD

	TRIFLOXYSTROBIN
	21
	10
	21
	N/A
	ALS IHD

	TRINEXAPAC-ETHYL
	22
	21
	21
	N/A
	ALS IHD


Interferences
HPLC-MS/MS is an extremely selective technique and interferences should only be encountered very rarely. Any interfering compounds would have to display the identical SRM transition at the same retention time, this is extremely unlikely in potable water samples. However, any compound, which passes through the extraction procedure, and has a similar liquid chromatographic retention time and mass spectrometric properties to the compound of interest, will cause interference. Samples containing high humic or fulvic loading have been demonstrated to not cause significant ion suppression for the compounds.
Performance of Method:

Range of Application:

Instrument WQQQ3, WQQQ4, WQQQ6 and WQQQ8
	Determinand
	Operational Calibration Range

	ATRAZINE
	LOQ - 0.150 µg L-1

	ATRAZINE-DESETHYL
	LOQ - 0.150 µg L-1

	ATRAZINE-DESISOPROPYL
	LOQ - 0.150 µg L-1

	AZINPHOS-METHYL
	LOQ - 0.150 µg L-1

	AZOXYSTROBIN
	LOQ - 0.150 µg L-1

	BENZOVINDIFLUPYR
	LOQ - 0.150 µg L-1

	BIXAFEN
	LOQ - 0.150 µg L-1

	BOSCALID
	LOQ - 0.150 µg L-1

	BROMACIL
	LOQ - 0.150 µg L-1

	CARBENDAZIM
	LOQ - 0.150 µg L-1

	CARBETAMIDE
	LOQ - 0.150 µg L-1

	CHLORFENVINPHOS
	LOQ - 0.150 µg L-1

	CHLORTOLURON
	LOQ - 0.150 µg L-1

	CYANAZINE
	LOQ - 0.150 µg L-1

	CYPROCONAZOLE
	LOQ - 0.150 µg L-1

	DICHLORVOS
	LOQ - 0.150 µg L-1

	DICLOFENAC
	LOQ - 0.150 µg L-1

	DIFENOCONAZOLE
	LOQ - 0.150 µg L-1

	DIMETHENAMID-P
	LOQ - 0.150 µg L-1

	DIMETHOATE
	LOQ - 0.150 µg L-1

	DIURON
	LOQ - 0.150 µg L-1

	EPOXICONAZOLE
	LOQ - 0.150 µg L-1

	EPTC
	LOQ - 0.150 µg L-1

	FLUFENACET
	LOQ - 0.150 µg L-1

	FLUOPYRAM
	LOQ - 0.150 µg L-1

	FLURTAMONE
	LOQ - 0.150 µg L-1

	FLUSILAZOLE
	LOQ - 0.150 µg L-1

	FLUTRIFOL
	LOQ - 0.150 µg L-1

	FLUXAPYROXAD
	LOQ - 0.150 µg L-1

	ISOPROTURON
	LOQ - 0.150 µg L-1

	LENACIL
	LOQ - 0.150 µg L-1

	LINURON
	LOQ - 0.150 µg L-1

	MALATHION
	LOQ - 0.150 µg L-1

	MEFENTRIFLUCONAZOLE
	LOQ - 0.150 µg L-1

	METAMITRON
	LOQ - 0.150 µg L-1

	METAZACHLOR
	LOQ - 0.150 µg L-1

	METCONAZOLE
	LOQ - 0.150 µg L-1

	METHABENZTHIAZURON
	LOQ - 0.150 µg L-1

	METHIOCARB
	LOQ - 0.150 µg L-1

	METOBROMURON
	LOQ - 0.150 µg L-1

	METOXURON
	LOQ - 0.150 µg L-1

	METRIBUZIN
	LOQ - 0.150 µg L-1

	METSULFURON-METHYL
	LOQ - 0.150 µg L-1

	MEVINPHOS
	LOQ - 0.150 µg L-1

	MONURON
	LOQ - 0.150 µg L-1

	OXADIXYL
	LOQ - 0.150 µg L-1

	OXAMYL
	LOQ - 0.150 µg L-1

	PROCHLORAZ
	LOQ - 0.150 µg L-1

	PROMETRYN
	LOQ - 0.150 µg L-1

	PROPACHLOR
	LOQ - 0.150 µg L-1

	PROPAZINE
	LOQ - 0.150 µg L-1

	PROPETAMPHOS
	LOQ - 0.150 µg L-1

	PROPICONAZOLE
	LOQ - 0.150 µg L-1

	PROSULFOCARB
	LOQ - 0.150 µg L-1

	PROTHIOCONAZOLE-DESTHIO
	LOQ - 0.150 µg L-1

	PYRACLOSTROBIN
	LOQ - 0.150 µg L-1

	QUINMERAC
	LOQ - 0.150 µg L-1

	SIMAZINE
	LOQ - 0.150 µg L-1

	SPIROXAMINE
	LOQ - 0.150 µg L-1

	TEBUCONAZOLE
	LOQ - 0.150 µg L-1

	TERBUTHYLAZINE
	LOQ - 0.150 µg L-1

	TERBUTRYN
	LOQ - 0.150 µg L-1

	THIFENSULFURON-METHYL
	LOQ - 0.150 µg L-1

	TRIAZOPHOS
	LOQ - 0.150 µg L-1

	TRIETAZINE
	LOQ - 0.150 µg L-1

	TRIFLOXYSTROBIN
	LOQ - 0.150 µg L-1

	TRINEXAPAC-ETHYL
	LOQ - 0.150 µg L-1


Limit of Quantification, Recoveries of Compounds, Bias and Uncertainty of Measurement:

Instrument 1, WQQQ3:
	Determinand
	UoM
	Direct Standards
	Elvington Treated Water 

(Hard Hardness)

	
	
	Low Standard
	HighStandard
	PCV Spike

	
	
	Recovery
	RSD
	Recovery
	RSD
	Recovery
	RSD

	Atrazine
	± 6.55 %
	98.9%
	1.3%
	100.2%
	1.0%
	98.6%
	1.2%

	Atrazine-desethyl
	± 23.32 %
	99.5%
	2.8%
	100.0%
	1.9%
	88.9%
	2.6%

	Atrazine-desisopropyl
	± 20.64 %
	98.8%
	3.7%
	100.2%
	2.5%
	90.3%
	2.4%

	Azinphos-methyl
	± 5.93 %
	99.3%
	1.8%
	99.7%
	1.2%
	99.9%
	1.2%

	Azoxystrobin
	± 10.08 %
	100.4%
	4.6%
	101.3%
	3.1%
	101.4%
	3.8%

	Benzovindiflupyr
	± 21.40%
	110.8%
	14.3%
	103.7%
	5.3%
	108.0%
	8.7%

	Bixafen
	± 11.17 %
	104.2%
	8.7%
	100.3%
	4.3%
	101.6%
	4.0%

	Boscalid
	± 8.38 %
	101.0%
	5.3%
	100.7%
	3.0%
	99.4%
	2.4%

	Bromacil
	± 9.27 %
	99.4%
	6.4%
	101.2%
	2.0%
	102.8%
	3.5%

	Carbendazim
	± 8.07 %
	98.1%
	1.7%
	100.1%
	1.3%
	98.8%
	1.1%

	Carbetamide
	± 8.68 %
	101.6%
	4.5%
	101.3%
	2.4%
	102.2%
	2.3%

	Chlorfenvinphos
	± 13.75 %
	101.4%
	5.7%
	101.5%
	5.9%
	104.3%
	5.2%

	Chlortoluron
	± 7.67 %
	100.1%
	2.7%
	99.9%
	3.0%
	99.4%
	2.4%

	Cyanazine
	± 6.38 %
	99.1%
	2.7%
	100.3%
	1.6%
	100.5%
	2.0%

	Cyproconazole
	± 8.62 %
	101.3%
	3.5%
	99.7%
	3.5%
	100.6%
	3.2%

	Dichlorvos
	± 7.80 %
	99.5%
	3.6%
	101.6%
	2.2%
	100.7%
	2.6%

	Diclofenac
	± 20.19%
	103.6%
	12.9%
	99.1%
	7.1%
	102.5%
	6.2%

	Difenoconazole
	± 10.10 %
	98.6%
	5.5%
	101.1%
	3.2%
	100.7%
	4.6%

	Dimethenamid-P
	± 10.27%
	101.2%
	5.2%
	103.3%
	3.1%
	103.5%
	3.2%

	Dimethoate
	± 6.94 %
	99.9%
	3.5%
	100.2%
	2.0%
	100.1%
	1.8%

	Diuron
	± 9.77 %
	101.9%
	5.0%
	102.7%
	2.0%
	102.3%
	2.9%

	Epoxiconazole
	± 7.70 %
	98.4%
	3.2%
	99.8%
	2.7%
	100.3%
	3.2%

	EPTC
	± 10.56 %
	100.4%
	7.8%
	100.2%
	2.8%
	101.4%
	3.6%

	Flufenacet
	± 13.31 %
	100.6%
	5.1%
	102.5%
	5.3%
	102.7%
	4.8%

	Fluopyram
	± 9.67 %
	102.2%
	5.9%
	100.9%
	3.3%
	101.2%
	2.8%

	Flurtamone
	± 7.15 %
	99.2%
	2.9%
	100.6%
	1.7%
	99.7%
	2.0%

	Flusilazole
	± 8.27 %
	99.8%
	3.0%
	100.2%
	3.7%
	101.2%
	2.8%

	Flutriafol
	± 9.50 %
	98.9%
	1.9%
	100.5%
	1.2%
	103.5%
	1.1%

	Fluxapyroxad
	± 10.10 %
	103.3%
	6.5%
	100.5%
	5.0%
	100.8%
	4.5%

	Isoproturon
	± 6.77 %
	99.3%
	2.5%
	100.4%
	1.5%
	100.1%
	1.6%

	Lenacil
	± 10.88 %
	99.0%
	4.1%
	100.9%
	2.8%
	103.6%
	2.5%

	Linuron
	± 10.43 %
	100.5%
	5.7%
	101.5%
	3.0%
	102.6%
	4.6%

	Malathion
	± 9.75 %
	101.5%
	3.9%
	100.0%
	3.0%
	100.5%
	3.0%

	Mefentrifluconazole
	± 14.85 %
	95.3%
	4.2%
	101.3%
	4.0%
	100.3%
	5.5%

	Metamitron
	± 21.67 %
	103.5%
	3.2%
	105.3%
	2.1%
	103.2%
	2.6%

	Metazachlor
	± 6.67 %
	100.8%
	3.4%
	101.5%
	1.7%
	100.5%
	2.0%

	Metconazole
	± 7.71 %
	98.6%
	2.4%
	99.8%
	1.8%
	100.3%
	2.4%

	Methabenzthiazuron
	± 7.16 %
	99.6%
	3.5%
	99.6%
	1.9%
	99.8%
	2.5%

	Methiocarb
	± 12.87 %
	99.8%
	4.2%
	102.2%
	2.9%
	104.3%
	2.5%

	Metobromuron
	± 10.77 %
	99.7%
	5.4%
	100.5%
	3.1%
	100.1%
	3.4%

	Metoxuron
	± 8.52 %
	99.8%
	3.7%
	99.9%
	3.4%
	101.3%
	2.9%

	Metribuzin
	± 6.74 %
	100.9%
	2.1%
	100.2%
	1.2%
	99.0%
	1.7%

	Metsulfuron-methyl
	± 15.26 %
	108.9%
	17.4%
	100.4%
	4.0%
	95.5%
	4.9%

	Mevinphos
	± 7.44 %
	99.7%
	4.9%
	101.3%
	2.0%
	100.5%
	2.1%

	Monuron
	± 6.84 %
	98.8%
	3.7%
	100.1%
	3.0%
	100.1%
	2.8%

	Oxadixyl
	± 8.66 %
	100.8%
	4.4%
	102.2%
	2.4%
	102.1%
	3.2%

	Oxamyl
	± 5.75 %
	97.8%
	3.1%
	98.7%
	2.0%
	98.6%
	1.2%

	Prochloraz
	± 12.52 %
	99.5%
	2.3%
	101.3%
	1.6%
	105.3%
	1.8%

	Prometryn
	± 5.92 %
	99.8%
	2.9%
	99.8%
	0.8%
	99.1%
	1.2%

	Propachlor
	± 7.06 %
	100.9%
	3.8%
	102.0%
	2.6%
	100.5%
	2.9%

	Propazine
	± 8.03 %
	99.7%
	1.9%
	100.4%
	1.2%
	97.4%
	1.6%

	Propetamphos
	± 11.08 %
	103.5%
	7.8%
	100.6%
	3.6%
	101.4%
	4.3%

	Propiconazole
	± 8.38 %
	99.9%
	4.5%
	100.6%
	2.6%
	101.1%
	2.9%

	Prosulfocarb
	± 17.62 %
	97.7%
	5.5%
	99.0%
	6.7%
	101.9%
	4.7%

	Prothioconazole-desthio
	± 8.48 %
	98.5%
	4.2%
	99.1%
	4.0%
	99.0%
	3.4%

	Pyraclostrobin
	± 12.78 %
	96.9%
	7.3%
	101.9%
	4.1%
	100.2%
	4.8%

	Quinmerac
	± 8.23 %
	99.4%
	3.3%
	99.2%
	2.3%
	98.8%
	4.0%

	Simazine
	± 5.94 %
	98.6%
	2.1%
	100.0%
	1.1%
	99.5%
	1.6%

	Spiroxamine
	± 9.51 %
	98.7%
	3.7%
	101.4%
	4.9%
	101.8%
	3.5%

	Tebuconazole
	± 6.69 %
	99.2%
	2.3%
	99.7%
	2.3%
	99.8%
	2.6%

	Terbuthylazine
	± 7.29 %
	100.0%
	1.6%
	100.4%
	1.2%
	98.0%
	2.0%

	Terbutryn
	± 5.76 %
	99.2%
	1.4%
	100.1%
	0.8%
	99.7%
	1.0%

	Thifensulfuron-methyl
	± 20.22 %
	100.3%
	5.5%
	99.9%
	3.8%
	103.8%
	4.0%

	Triazophos
	± 9.89 %
	100.1%
	5.4%
	99.3%
	3.1%
	99.7%
	3.7%

	Trietazine
	± 10.35 %
	97.7%
	2.6%
	99.9%
	1.7%
	103.9%
	1.7%

	Trifloxystrobin
	± 34.32 %
	104.0%
	11.3%
	101.5%
	7.5%
	101.7%
	7.8%

	Trinexapac-ethyl
	± 14.28 %
	102.3%
	7.9%
	98.6%
	5.7%
	102.1%
	5.0%


Instrument 2, WQQQ4:
	Determinand
	UoM
	Direct Standards
	Elvington Treated Water 

(Hard Hardness)

	
	
	Low Standard
	HighStandard
	PCV Spike

	
	
	Recovery
	RSD
	Recovery
	RSD
	Recovery
	RSD

	Atrazine
	± 6.32 %
	99.9%
	3.3%
	99.7%
	1.5%
	99.9%
	1.6%

	Atrazine-desethyl
	± 8.93 %
	102.1%
	2.7%
	100.4%
	3.0%
	99.2%
	3.3%

	Atrazine-desisopropyl
	± 21.78 %
	100.6%
	4.0%
	98.1%
	3.5%
	90.5%
	5.3%

	Azinphos-methyl
	± 6.89 %
	101.1%
	1.7%
	99.8%
	1.3%
	100.8%
	1.7%

	Azoxystrobin
	± 18.53 %
	96.4%
	8.6%
	98.4%
	7.4%
	100.0%
	8.8%

	Benzovindiflupyr
	± 27.50 %
	103.6%
	7.8%
	102.2%
	8.6%
	106.9%
	6.5%

	Bixafen
	± 10.61 %
	100.7%
	4.5%
	101.5%
	5.7%
	100.7%
	3.7%

	Boscalid
	± 10.37 %
	98.0%
	5.1%
	99.0%
	4.7%
	98.8%
	4.9%

	Bromacil
	± 18.54 %
	96.8%
	6.8%
	97.7%
	6.6%
	99.1%
	6.5%

	Carbendazim
	± 9.78 %
	99.3%
	1.5%
	100.0%
	0.7%
	100.9%
	1.3%

	Carbetamide
	± 6.82 %
	102.8%
	3.7%
	99.4%
	2.1%
	100.1%
	2.2%

	Chlorfenvinphos
	± 12.84 %
	97.9%
	4.6%
	100.4%
	5.7%
	102.9%
	3.8%

	Chlortoluron
	± 10.03 %
	103.1%
	3.9%
	99.9%
	3.7%
	99.2%
	5.0%

	Cyanazine
	± 6.79 %
	99.7%
	1.9%
	99.2%
	1.6%
	99.8%
	1.9%

	Cyproconazole
	± 9.09 %
	100.5%
	3.3%
	100.6%
	4.1%
	100.8%
	3.1%

	Dichlorvos
	± 8.34 %
	100.5%
	3.2%
	99.2%
	2.1%
	100.4%
	2.2%

	Diclofenac
	± 17.95 %
	97.9%
	4.8%
	100.3%
	5.1%
	100.5%
	5.6%

	Difenoconazole
	± 11.11 %
	100.9%
	3.0%
	100.4%
	2.6%
	103.8%
	2.3%

	Dimethenamid-P
	± 20.63 %
	101.8%
	6.3%
	102.3%
	6.1%
	103.6%
	6.1%

	Dimethoate
	± 11.03 %
	100.8%
	2.8%
	100.4%
	2.1%
	102.2%
	2.5%

	Diuron
	± 13.98 %
	102.9%
	6.9%
	99.5%
	5.8%
	102.3%
	6.0%

	Epoxiconazole
	± 8.53 %
	99.9%
	2.6%
	100.0%
	3.5%
	101.0%
	2.4%

	EPTC
	± 9.81 %
	100.8%
	4.9%
	100.3%
	2.8%
	102.2%
	3.3%

	Flufenacet
	± 23.32 %
	96.8%
	10.0%
	95.6%
	9.1%
	96.3%
	10.0%

	Fluopyram
	± 15.55 %
	99.0%
	4.1%
	99.9%
	3.8%
	99.8%
	6.3%

	Flurtamone
	± 7.19 %
	99.3%
	3.2%
	98.6%
	2.7%
	99.7%
	3.6%

	Flusilazole
	± 9.94 %
	99.4%
	3.2%
	99.8%
	3.2%
	102.7%
	2.4%

	Flutriafol
	± 10.32 %
	100.3%
	3.3%
	99.5%
	2.5%
	103.0%
	2.2%

	Fluxapyroxad
	± 10.46 %
	101.3%
	3.3%
	101.1%
	6.8%
	99.6%
	3.0%

	Isoproturon
	± 9.57 %
	101.8%
	4.3%
	101.3%
	3.7%
	102.0%
	3.8%

	Lenacil
	± 11.63 %
	98.2%
	6.3%
	98.8%
	4.3%
	102.3%
	4.7%

	Linuron
	± 13.38 %
	99.6%
	5.5%
	100.4%
	5.4%
	98.7%
	6.2%

	Malathion
	± 8.78 %
	101.6%
	3.6%
	99.8%
	2.0%
	101.6%
	2.1%

	Mefentrifluconazole
	± 10.98%
	94.8%
	4.01
	99.7%
	2.56%
	101.2%
	4.9%

	Metamitron
	± 12.94 %
	100.9%
	5.2%
	103.2%
	5.5%
	96.4%
	7.3%

	Metazachlor
	± 10.38 %
	98.2%
	5.1%
	98.6%
	3.3%
	100.8%
	4.9%

	Metconazole
	± 9.32 %
	98.2%
	3.6%
	99.5%
	3.7%
	99.5%
	3.2%

	Methabenzthiazuron
	± 14.23 %
	105.0%
	3.9%
	103.9%
	3.6%
	105.1%
	3.7%

	Methiocarb
	± 15.37 %
	102.3%
	7.0%
	101.6%
	5.6%
	102.2%
	8.2%

	Metobromuron
	± 13.98 %
	98.0%
	6.0%
	98.3%
	5.5%
	97.9%
	5.9%

	Metoxuron
	± 19.59 %
	101.5%
	4.7%
	100.2%
	3.2%
	108.0%
	3.1%

	Metribuzin
	± 10.10 %
	99.9%
	2.0%
	100.7%
	1.9%
	102.8%
	2.3%

	Metsulfuron-methyl
	± 19.03 %
	101.3%
	8.9%
	96.7%
	3.8%
	99.1%
	4.6%

	Mevinphos
	± 9.03 %
	101.0%
	4.1%
	99.9%
	3.3%
	102.3%
	2.7%

	Monuron
	± 7.28 %
	101.1%
	2.7%
	100.4%
	2.6%
	100.9%
	2.0%

	Oxadixyl
	± 8.99 %
	96.6%
	4.5%
	99.0%
	5.0%
	102.0%
	3.2%

	Oxamyl
	± 5.32 %
	98.0%
	2.5%
	100.1%
	1.5%
	100.0%
	1.2%

	Prochloraz
	± 8.96 %
	100.9%
	3.7%
	101.0%
	2.2%
	101.8%
	3.1%

	Prometryn
	± 7.09 %
	99.5%
	1.8%
	100.3%
	1.5%
	100.9%
	1.1%

	Propachlor
	± 13.55 %
	97.8%
	7.4%
	97.4%
	4.5%
	98.0%
	5.8%

	Propazine
	± 6.88 %
	99.8%
	3.3%
	99.7%
	1.4%
	99.7%
	1.5%

	Propetamphos
	± 9.40 %
	102.1%
	5.7%
	98.6%
	2.5%
	100.5%
	5.0%

	Propiconazole
	± 6.98 %
	99.2%
	2.2%
	98.5%
	2.3%
	99.7%
	1.8%

	Prosulfocarb
	± 22.63 %
	94.9%
	10.0%
	96.6%
	6.5%
	94.9%
	8.6%

	Prothioconazole-desthio
	± 11.14 %
	98.0%
	4.7%
	99.7%
	5.7%
	99.4%
	4.0%

	Pyraclostrobin
	± 7.40 %
	98.3%
	2.5%
	99.2%
	2.4%
	99.6%
	2.4%

	Quinmerac
	± 30.88 %
	102.8%
	5.4%
	100.9%
	3.6%
	114.3%
	7.2%

	Simazine
	± 6.85 %
	100.4%
	2.8%
	99.5%
	1.4%
	100.3%
	1.7%

	Spiroxamine
	± 14.09 %
	100.0%
	5.0%
	99.1%
	5.7%
	104.8%
	5.5%

	Tebuconazole
	± 8.58 %
	99.2%
	2.8%
	100.0%
	2.1%
	100.9%
	2.3%

	Terbuthylazine
	± 7.27 %
	100.4%
	2.2%
	100.4%
	1.8%
	101.1%
	1.6%

	Terbutryn
	± 7.66 %
	100.1%
	1.7%
	100.2%
	1.3%
	100.8%
	1.3%

	Thifensulfuron-methyl
	± 14.15 %
	95.4%
	6.3%
	97.9%
	4.8%
	103.1%
	4.0%

	Triazophos
	± 10.69 %
	101.5%
	4.5%
	101.0%
	3.0%
	102.9%
	2.2%

	Trietazine
	± 7.38 %
	99.7%
	2.7%
	100.0%
	3.4%
	101.5%
	1.9%

	Trifloxystrobin
	± 21.58 %
	99.6%
	10.0%
	98.7%
	7.9%
	101.2%
	6.2%

	Trinexapac-ethyl
	± 10.06 %
	104.4%
	6.1%
	101.4%
	6.2%
	102.2%
	3.3%


Instrument 3, WQQQ6:
	Determinand
	UoM
	Direct Standards
	Elvington Treated Water 

(Hard Hardness)

	
	
	Low Standard
	HighStandard
	PCV Spike

	
	
	Recovery
	RSD
	Recovery
	RSD
	Recovery
	RSD

	Atrazine
	6.408%
	97.06%
	1.48%
	99.83%
	1.04%
	98.32%
	1.70%

	Atrazine-desethyl
	19.769%
	101.65%
	9.04%
	100.21%
	2.47%
	94.02%
	8.93%

	Atrazine-desisopropyl
	36.463%
	97.45%
	4.70%
	98.08%
	4.60%
	82.03%
	9.59%

	Azinphos-methyl
	6.513%
	100.62%
	2.19%
	100.60%
	1.42%
	101.47%
	1.78%

	Azoxystrobin
	7.793%
	94.75%
	4.56%
	100.08%
	2.15%
	97.60%
	3.63%

	Benzovindiflupyr
	18.479%
	96.77%
	8.29%
	100.22%
	6.39%
	98.26%
	7.57%

	Bixafen
	12.488%
	98.70%
	6.55%
	100.42%
	3.13%
	100.80%
	5.00%

	Boscalid
	8.993%
	98.02%
	4.85%
	100.44%
	3.02%
	96.63%
	2.99%

	Bromacil
	13.735%
	97.28%
	5.18%
	99.66%
	2.21%
	96.85%
	3.74%

	Carbendazim
	5.278%
	95.73%
	1.36%
	100.13%
	1.20%
	99.18%
	1.07%

	Carbetamide
	12.416%
	97.69%
	3.44%
	100.26%
	2.59%
	99.69%
	3.91%

	Chlorfenvinphos
	16.221%
	97.53%
	8.51%
	100.46%
	3.10%
	100.64%
	5.55%

	Chlortoluron
	10.338%
	99.00%
	3.42%
	100.93%
	2.77%
	100.36%
	4.27%

	Cyanazine
	6.737%
	95.72%
	2.06%
	99.82%
	1.70%
	99.15%
	2.57%

	Cyproconazole
	9.564%
	96.74%
	3.50%
	99.66%
	2.64%
	98.70%
	2.44%

	Dichlorvos
	9.387%
	98.13%
	3.88%
	100.67%
	3.10%
	100.00%
	2.64%

	Diclofenac
	21.689%
	96.21%
	9.96%
	98.86%
	8.80%
	101.42%
	6.14%

	Difenoconazole
	14.359%
	93.98%
	5.01%
	98.21%
	4.89%
	102.39%
	4.27%

	Dimethenamid-P
	10.896%
	101.06%
	4.54%
	100.89%
	3.85%
	95.37%
	6.60%

	Dimethoate
	12.393%
	95.93%
	4.50%
	100.10%
	2.67%
	101.36%
	5.74%

	Diuron
	10.074%
	106.91%
	4.50%
	100.35%
	3.49%
	100.28%
	3.52%

	Epoxiconazole
	11.551%
	94.81%
	4.97%
	100.06%
	3.71%
	101.50%
	3.94%

	EPTC
	15.069%
	100.11%
	5.66%
	99.96%
	4.77%
	103.09%
	4.39%

	Flufenacet
	17.417%
	100.85%
	7.92%
	102.37%
	7.30%
	99.28%
	6.91%

	Fluopyram
	6.654%
	96.22%
	2.71%
	99.56%
	2.18%
	98.75%
	2.26%

	Flurtamone
	6.844%
	98.41%
	2.80%
	100.61%
	1.55%
	97.77%
	2.19%

	Flusilazole
	15.084%
	88.31%
	4.72%
	99.05%
	3.33%
	97.63%
	3.29%

	Flutriafol
	9.184%
	94.39%
	2.39%
	99.56%
	1.95%
	101.90%
	2.38%

	Fluxapyroxad
	14.892%
	102.33%
	4.11%
	100.18%
	3.39%
	97.80%
	4.40%

	Isoproturon
	6.909%
	100.17%
	2.57%
	100.04%
	1.83%
	100.56%
	1.73%

	Lenacil
	9.758%
	92.75%
	3.51%
	99.43%
	2.02%
	98.93%
	1.61%

	Linuron
	10.295%
	98.19%
	5.32%
	100.54%
	2.28%
	100.43%
	2.52%

	Malathion
	8.845%
	99.33%
	3.23%
	100.07%
	2.94%
	100.15%
	3.06%

	Mefentrifluconazole
	10.336%
	93.78%
	5.09%
	99.20%
	3.15%
	99.81%
	3.71%

	Metamitron
	16.475%
	95.45%
	3.94%
	99.72%
	4.15%
	98.67%
	6.25%

	Metazachlor
	7.539%
	100.11%
	3.64%
	100.62%
	1.49%
	97.60%
	3.31%

	Metconazole
	10.094%
	95.13%
	2.94%
	100.24%
	2.37%
	98.27%
	1.94%

	Methabenzthiazuron
	9.993%
	92.00%
	2.82%
	99.42%
	2.42%
	99.77%
	2.73%

	Methiocarb
	16.958%
	98.19%
	4.91%
	100.47%
	4.41%
	106.46%
	5.54%

	Metobromuron
	14.92%
	99.86%
	5.60%
	102.06%
	4.44%
	94.90%
	3.89%

	Metoxuron
	9.409%
	97.74%
	2.56%
	98.99%
	2.66%
	105.05%
	2.74%

	Metribuzin
	14.286%
	95.51%
	3.66%
	99.89%
	2.99%
	101.51%
	4.76%

	Metsulfuron-methyl
	17.828%
	96.38%
	11.61%
	101.38%
	3.68%
	100.07%
	5.27%

	Mevinphos
	13.224%
	96.12%
	3.69%
	99.14%
	3.36%
	103.72%
	2.52%

	Monuron
	10.426%
	101.01%
	4.23%
	101.27%
	3.53%
	101.56%
	3.09%

	Oxadixyl
	11.819%
	96.96%
	4.16%
	99.62%
	3.84%
	102.73%
	2.74%

	Oxamyl
	4.643%
	96.98%
	1.19%
	99.95%
	1.05%
	99.86%
	1.02%

	Prochloraz
	11.939%
	92.16%
	4.62%
	100.52%
	3.25%
	99.46%
	3.48%

	Prometryn
	5.58%
	95.86%
	2.08%
	100.01%
	1.36%
	98.51%
	1.28%

	Propachlor
	10.474%
	101.40%
	3.47%
	100.38%
	2.15%
	94.80%
	4.83%

	Propazine
	6.97%
	96.13%
	1.87%
	99.85%
	1.33%
	101.91%
	2.14%

	Propetamphos
	14.008%
	102.98%
	6.63%
	101.45%
	3.92%
	97.77%
	4.77%

	Propiconazole
	8.62%
	94.46%
	4.44%
	99.22%
	3.26%
	99.81%
	2.48%

	Prosulfocarb
	12.668%
	100.66%
	6.11%
	100.71%
	3.41%
	99.98%
	3.75%

	Prothioconazole-desthio
	10.018%
	95.95%
	4.36%
	100.16%
	3.94%
	99.95%
	3.62%

	Pyraclostrobin
	18.664%
	96.04%
	6.02%
	99.47%
	5.75%
	99.03%
	7.10%

	Quinmerac
	6.607%
	95.11%
	2.96%
	100.25%
	2.02%
	101.19%
	1.74%

	Simazine
	8.393%
	96.25%
	1.83%
	99.82%
	1.57%
	99.56%
	1.73%

	Spiroxamine
	10.417%
	94.21%
	4.31%
	100.69%
	3.88%
	100.98%
	3.08%

	Tebuconazole
	7.909%
	95.81%
	2.24%
	99.60%
	2.73%
	98.32%
	2.08%

	Terbuthylazine
	9.379%
	97.83%
	1.69%
	100.56%
	1.38%
	103.06%
	2.87%

	Terbutryn
	4.752%
	96.17%
	1.40%
	100.01%
	1.07%
	99.33%
	1.20%

	Thifensulfuron-methyl
	21.343%
	99.28%
	5.59%
	100.69%
	3.40%
	104.94%
	4.42%

	Triazophos
	12.584%
	97.80%
	4.84%
	100.65%
	3.26%
	104.52%
	4.98%

	Trietazine
	15.575%
	95.46%
	1.65%
	99.60%
	1.56%
	105.77%
	4.53%

	Trifloxystrobin
	11.363%
	95.00%
	4.50%
	98.29%
	2.74%
	100.09%
	4.78%

	Trinexapac-ethyl
	10.15%
	95.76%
	7.34%
	101.21%
	2.86%
	101.43%
	3.75%


Instrument 4: WQQQ8

	Determinand
	UoM
	Direct Standards
	Elvington Treated Water 

(Hard Hardness)

	
	
	Low Standard
	HighStandard
	PCV Spike

	
	
	Recovery
	RSD
	Recovery
	RSD
	Recovery
	RSD

	Atrazine
	± 5.94 %
	99.9%
	1.0%
	99.8%
	0.8%
	99.0%
	1.3%

	Atrazine-desethyl
	± 9.85 %
	100.8%
	2.1%
	99.9%
	0.7%
	96.1%
	2.0%

	Atrazine-desisopropyl
	± 24.21 %
	100.2%
	2.3%
	99.2%
	2.6%
	88.5%
	4.8%

	Azinphos-methyl
	± 7.21 %
	102.2%
	2.1%
	99.6%
	1.5%
	101.2%
	2.0%

	Azoxystrobin
	± 11.58 %
	99.0%
	4.0%
	100.2%
	4.0%
	101.0%
	4.0%

	Benzovindiflupyr
	± 15.78 %
	99.2%
	8.3%
	99.8%
	7.3%
	102.8%
	7.5%

	Bixafen
	± 9.82 %
	100.7%
	5.3%
	99.2%
	2.6%
	100.0%
	4.6%

	Boscalid
	± 7.09 %
	98.5%
	3.1%
	100.0%
	1.4%
	99.6%
	2.9%

	Bromacil
	± 9.90 %
	99.5%
	2.4%
	99.6%
	2.8%
	102.1%
	3.5%

	Carbendazim
	± 6.03 %
	100.6%
	1.5%
	100.0%
	1.4%
	100.2%
	1.8%

	Carbetamide
	± 6.71 %
	100.4%
	3.4%
	100.0%
	1.7%
	100.9%
	1.7%

	Chlorfenvinphos
	± 11.50 %
	99.3%
	5.1%
	97.8%
	4.4%
	97.9%
	5.3%

	Chlortoluron
	± 6.21 %
	99.0%
	1.7%
	99.8%
	1.2%
	99.7%
	1.4%

	Cyanazine
	± 5.85 %
	98.7%
	1.8%
	99.7%
	1.0%
	99.3%
	1.3%

	Cyproconazole
	± 7.27 %
	99.7%
	2.1%
	100.7%
	1.9%
	100.8%
	2.2%

	Dichlorvos
	± 6.33 %
	99.3%
	2.9%
	99.5%
	1.7%
	100.4%
	1.6%

	Diclofenac
	± 25.14 %
	105.5%
	7.2%
	102.3%
	7.9%
	103.0%
	9.8%

	Difenoconazole
	± 8.36 %
	98.7%
	3.8%
	99.5%
	3.0%
	100.2%
	3.6%

	Dimethenamid-P
	± 9.54 %
	100.9%
	3.4%
	101.1%
	2.7%
	102.8%
	2.7%

	Dimethoate
	± 6.65 %
	99.7%
	2.1%
	100.1%
	2.2%
	101.2%
	2.0%

	Diuron
	± 7.73 %
	102.1%
	2.4%
	100.4%
	2.7%
	101.6%
	2.7%

	Epoxiconazole
	± 6.65 %
	99.5%
	1.7%
	99.9%
	1.8%
	100.1%
	1.8%

	EPTC
	± 6.58 %
	98.1%
	1.8%
	98.8%
	1.5%
	99.2%
	2.2%

	Flufenacet
	± 11.22 %
	99.4%
	4.7%
	100.1%
	4.1%
	100.4%
	5.3%

	Fluopyram
	± 8.32 %
	100.2%
	3.5%
	99.6%
	2.9%
	100.1%
	3.4%

	Flurtamone
	± 7.03 %
	100.6%
	2.4%
	99.9%
	1.3%
	100.9%
	2.2%

	Flusilazole
	± 7.66 %
	99.7%
	2.7%
	99.7%
	2.0%
	99.8%
	2.3%

	Flutriafol
	± 6.39 %
	99.4%
	1.0%
	100.2%
	0.6%
	101.4%
	2.6%

	Fluxapyroxad
	± 9.70 %
	101.7%
	4.4%
	100.0%
	3.4%
	100.3%
	3.8%

	Isoproturon
	± 6.06 %
	101.3%
	1.0%
	99.8%
	1.2%
	100.7%
	1.7%

	Lenacil
	± 7.48 %
	98.8%
	2.9%
	100.0%
	1.7%
	102.0%
	2.8%

	Linuron
	± 8.52 %
	100.8%
	3.2%
	100.7%
	2.6%
	99.5%
	3.3%

	Malathion
	± 7.80 %
	99.0%
	2.5%
	99.3%
	1.4%
	100.0%
	5.3%

	Mefentrifluconazole
	± 11.93 %
	95.2%
	4.0%
	99.2%
	3.9%
	102.2%
	4.7%

	Metamitron
	± 13.06 %
	101.8%
	3.9%
	104.8%
	4.9%
	105.0%
	5.8%

	Metazachlor
	± 6.57 %
	100.5%
	1.9%
	100.0%
	1.1%
	101.1%
	2.3%

	Metconazole
	± 5.79 %
	99.6%
	1.2%
	99.9%
	1.2%
	99.6%
	1.7%

	Methabenzthiazuron
	± 5.99 %
	100.0%
	1.6%
	99.8%
	1.3%
	99.3%
	1.4%

	Methiocarb
	± 9.49 %
	101.3%
	4.2%
	100.9%
	3.1%
	101.7%
	3.3%

	Metobromuron
	± 8.35 %
	101.1%
	5.1%
	101.4%
	3.2%
	101.8%
	3.1%

	Metoxuron
	± 7.50 %
	98.9%
	1.5%
	99.5%
	1.3%
	102.3%
	1.6%

	Metribuzin
	± 7.17 %
	100.9%
	2.2%
	101.1%
	1.4%
	102.1%
	1.9%

	Metsulfuron-methyl
	± 11.35 %
	101.7%
	5.7%
	99.2%
	2.6%
	97.9%
	4.8%

	Mevinphos
	± 6.76 %
	99.6%
	2.0%
	99.5%
	1.7%
	101.3%
	1.4%

	Monuron
	± 6.01 %
	99.0%
	1.8%
	99.8%
	1.1%
	99.4%
	1.7%

	Oxadixyl
	± 6.52 %
	98.3%
	2.8%
	99.3%
	2.0%
	99.6%
	2.2%

	Oxamyl
	± 5.16 %
	97.5%
	2.6%
	99.0%
	1.1%
	98.8%
	1.0%

	Prochloraz
	± 14.01 %
	101.9%
	6.4%
	101.9%
	4.1%
	104.0%
	4.2%

	Prometryn
	± 5.99 %
	98.8%
	1.7%
	99.3%
	1.2%
	99.1%
	1.4%

	Propachlor
	± 6.79 %
	100.7%
	2.0%
	100.3%
	1.7%
	100.7%
	2.2%

	Propazine
	± 6.45 %
	100.1%
	1.5%
	99.6%
	1.3%
	98.8%
	1.5%

	Propetamphos
	± 9.41 %
	98.6%
	4.0%
	97.9%
	2.3%
	97.1%
	3.2%

	Propiconazole
	± 5.75 %
	99.4%
	1.4%
	100.1%
	1.3%
	99.9%
	1.2%

	Prosulfocarb
	± 13.91 %
	100.5%
	8.3%
	101.3%
	4.3%
	100.9%
	6.4%

	Prothioconazole-desthio
	± 7.93 %
	99.5%
	1.7%
	100.2%
	2.4%
	100.5%
	2.9%

	Pyraclostrobin
	± 13.58 %
	102.9%
	5.0%
	102.1%
	5.0%
	102.0%
	5.2%

	Quinmerac
	± 22.40 %
	100.3%
	1.7%
	99.4%
	1.5%
	110.6%
	6.4%

	Simazine
	± 5.99 %
	99.4%
	1.3%
	99.6%
	0.9%
	99.9%
	1.5%

	Spiroxamine
	± 8.77 %
	100.9%
	3.8%
	98.4%
	2.0%
	100.1%
	3.9%

	Tebuconazole
	± 6.10 %
	100.1%
	1.6%
	100.6%
	1.2%
	100.9%
	1.1%

	Terbuthylazine
	± 6.50 %
	100.1%
	1.3%
	99.6%
	1.2%
	99.0%
	1.6%

	Terbutryn
	± 5.93 %
	98.9%
	1.6%
	99.3%
	1.1%
	99.2%
	1.1%

	Thifensulfuron-methyl
	± 13.50 %
	99.8%
	3.9%
	99.7%
	3.3%
	105.2%
	5.7%

	Triazophos
	± 8.15 %
	99.0%
	3.3%
	98.9%
	2.9%
	101.1%
	3.5%

	Trietazine
	± 6.03 %
	98.5%
	1.5%
	99.6%
	0.9%
	100.1%
	2.2%

	Trifloxystrobin
	± 24.92 %
	105.0%
	8.7%
	102.7%
	10.4%
	101.8%
	12.4%

	Trinexapac-ethyl
	± 11.88 %
	100.4%
	3.0%
	99.7%
	4.0%
	103.1%
	4.5%


The highest LOQ has been applied across all instruments.

	Determinand
	WQQQ3

Limit of Quantification
(µg L-1)
	WQQQ4

Limit of Quantification
(µg L-1)
	WQQQ6

Limit of Quantification
(µg L-1)
	WQQQ8

Limit of Quantification
(µg L-1)
	METHOD (Standardised) LOQ

(µg L-1)

	Atrazine
	0.006
	0.004
	0.003
	0.004
	0.006

	Atrazine-desethyl
	0.005
	0.005
	0.004
	0.005
	0.005

	Atrazine-desisopropyl
	0.009
	0.008
	0.005
	0.006
	0.009

	Azinphos-methyl
	0.004
	0.004
	0.005
	0.005
	0.005

	Azoxystrobin
	0.004
	0.005
	0.005
	0.006
	0.006

	Benzovindiflupyr
	0.011
	0.012
	0.014
	0.013
	0.014

	Bixafen
	0.012
	0.011
	0.008
	0.009
	0.012

	Boscalid
	0.008
	0.007
	0.007
	0.007
	0.008

	Bromacil
	0.011
	0.010
	0.005
	0.007
	0.011

	Carbendazim
	0.004
	0.004
	0.004
	0.005
	0.005

	Carbetamide
	0.008
	0.009
	0.008
	0.004
	0.009

	Chlorfenvinphos
	0.013
	0.011
	0.007
	0.007
	0.013

	Chlortoluron
	0.005
	0.005
	0.004
	0.005
	0.005

	Cyanazine
	0.007
	0.007
	0.004
	0.005
	0.007

	Cyproconazole
	0.006
	0.006
	0.005
	0.006
	0.006

	Dichlorvos
	0.008
	0.007
	0.005
	0.005
	0.008

	Diclofenac
	0.019
	0.013
	0.015
	0.008
	0.019

	Difenoconazole
	0.009
	0.005
	0.008
	0.007
	0.009

	Dimethenamid-P
	0.006
	0.009
	0.008
	0.004
	0.009

	Dimethoate
	0.007
	0.007
	0.007
	0.006
	0.007

	Diuron
	0.009
	0.009
	0.007
	0.006
	0.009

	Epoxiconazole
	0.005
	0.005
	0.006
	0.004
	0.006

	EPTC
	0.011
	0.022
	0.010
	0.012
	0.022

	Flufenacet
	0.005
	0.008
	0.008
	0.007
	0.008

	Fluopyram
	0.004
	0.004
	0.005
	0.005
	0.005

	Flurtamone
	0.006
	0.007
	0.004
	0.005
	0.007

	Flusilazole
	0.007
	0.006
	0.009
	0.004
	0.009

	Flutriafol
	0.004
	0.005
	0.004
	0.004
	0.005

	Fluxapyroxad
	0.010
	0.010
	0.007
	0.007
	0.010

	Isoproturon
	0.004
	0.004
	0.004
	0.005
	0.005

	Lenacil
	0.006
	0.010
	0.004
	0.005
	0.010

	Linuron
	0.004
	0.009
	0.007
	0.007
	0.009

	Malathion
	0.009
	0.010
	0.006
	0.006
	0.010

	Mefentrifluconazole
	0.009
	0.006
	0.006
	0.004
	0.009

	Metamitron
	0.005
	0.007
	0.006
	0.008
	0.008

	Metazachlor
	0.005
	0.005
	0.004
	0.005
	0.005

	Metconazole
	0.006
	0.004
	0.003
	0.004
	0.006

	Methabenzthiazuron
	0.006
	0.003
	0.004
	0.005
	0.006

	Methiocarb
	0.006
	0.005
	0.006
	0.009
	0.009

	Metobromuron
	0.007
	0.010
	0.010
	0.007
	0.010

	Metoxuron
	0.006
	0.005
	0.004
	0.004
	0.006

	Metribuzin
	0.005
	0.005
	0.006
	0.006
	0.006

	Metsulfuron-methyl
	0.018
	0.016
	0.008
	0.009
	0.018

	Mevinphos
	0.011
	0.008
	0.009
	0.006
	0.011

	Monuron
	0.005
	0.005
	0.004
	0.005
	0.005

	Oxadixyl
	0.012
	0.010
	0.012
	0.007
	0.012

	Oxamyl
	0.003
	0.004
	0.004
	0.005
	0.005

	Prochloraz
	0.006
	0.007
	0.010
	0.007
	0.010

	Prometryn
	0.004
	0.005
	0.004
	0.004
	0.005

	Propachlor
	0.005
	0.007
	0.005
	0.004
	0.007

	Propazine
	0.004
	0.005
	0.003
	0.005
	0.005

	Propetamphos
	0.007
	0.009
	0.006
	0.007
	0.009

	Propiconazole
	0.006
	0.007
	0.009
	0.007
	0.009

	Prosulfocarb
	0.006
	0.005
	0.006
	0.004
	0.006

	Prothioconazole-desthio
	0.007
	0.005
	0.004
	0.006
	0.007

	Pyraclostrobin
	0.009
	0.008
	0.009
	0.007
	0.009

	Quinmerac
	0.004
	0.006
	0.005
	0.006
	0.006

	Simazine
	0.005
	0.005
	0.004
	0.004
	0.005

	Spiroxamine
	0.006
	0.009
	0.004
	0.003
	0.009

	Tebuconazole
	0.006
	0.003
	0.004
	0.004
	0.006

	Terbuthylazine
	0.005
	0.005
	0.004
	0.004
	0.005

	Terbutryn
	0.004
	0.004
	0.004
	0.004
	0.004

	Thifensulfuron-methyl
	0.009
	0.013
	0.008
	0.007
	0.013

	Triazophos
	0.009
	0.005
	0.006
	0.006
	0.009

	Trietazine
	0.004
	0.005
	0.003
	0.004
	0.005

	Trifloxystrobin
	0.007
	0.007
	0.006
	0.006
	0.007

	Trinexapac-ethyl
	0.009
	0.010
	0.005
	0.007
	0.010


References:

Determination of Phenyl Urea and Triazine Herbicides in Potable and Groundwater by LC/MS Using API-ESI Selective Ion Monitoring and Direct Large Volume Aqueous Injection, Agilent Technologies Application Note.

Agilent 6410 QQQ LC/MS Techniques and Operation, Agilent Technologies Course Number R1893A, Student Manuals Volumes 1 and 2.

Agilent 1200 Series, Reference Manuals

Agilent 6410 Triple Quad LC/MS, Maintenance and Familiarization Guides.

Transport of sulfonylureas from a barley field in Norway: Field and laboratory studies, Bioforsk Report Vol.3 Nr. 105 2008, ISBN-nr.:978-82-17-00409-7
	WPC71
	Last Updated: 19/03/26
	Page 14 of 14



[image: image1.jpg]