	METHOD STATEMENT
	[image: image1.jpg]N







Determinand:
Determination of Aluminium, Arsenic, Cadmium, Chromium, Cobalt, Copper, Iron, Manganese, Molybdenum, Nickel, Phosphorous, Lead, Selenium, Thallium, Vanadium and Zinc
Matrix:

Sample Type: final effluents, trade discharge, crude sewage and other similar waste water samples. Filtered samples may also be analysed using this method.

NOTE: For the Agilent 7900 ICPMS this method is not suitable for the analysis of Selenium in crude sewages.
MCerts accreditation is provided for pre-filtered on-site samples. Dissolved metals analysis performed in the lab is not covered by MCerts accreditation.
Principle of Method:

The method describes a technique for multi – element determination of trace metals in solution. The basis of the method is the measurement of ions produced by an inductively coupled plasma and detected using a mass spectrometer. 
The Agilent 7900 uses helium and high energy helium collision mode to reduce any potential interferences.
Acidified samples are nebulised and the aerosol that is produced is transported to the plasma torch where excitation of the metal atoms occurs. Excitation is due to the high temperatures (up to 6,000-K) produced by the radio frequency inductively coupled plasma. The metal ions thus produced pass through an interface region into the mass spectrometer. There the ions are separated by a quadrupole where only ions having a specific mass to charge ratio are passed through at any moment in time. The dual mode detector then detects these ions and the resulting electrical signals are processed into digital information that is used to indicate ion intensity and subsequently elemental concentration.
Sampling and Sample Preparation:

Samples are normally taken in 1L PET bottles.
Samples should be prepared within 7 days of sampling (Coventry in-house data)

Prepared samples are stable for one month. Cd, Cr, Cu, Pb, Ni, Zn and As are stable for 1 month. (ISO 5667 – 3:2024)

If sample preparation cannot be immediately undertaken, samples should be refrigerated at 3 ± 2°C.  Once digested, samples are stored at room temperature.
Interferences
Careful choice of plasma conditions, interference equations and isotopes are all designed to minimise interference.
Performance of Method:

Range of Application:

	Element
	Agilent 7900 µg/l

	Al
	LOD - 1000

	P
	LOD - 2500

	Cr
	LOD - 250

	Fe
	LOD - 2500

	Mn
	LOD - 1000

	Co
	LOD - 100

	Ni
	LOD - 250

	Cu
	LOD - 250

	Zn
	LOD - 250

	As
	LOD - 100

	Se
	LOD - 100

	Mo
	LOD - 100

	Cd
	LOD - 100

	Tl
	LOD - 100

	Pb
	LOD - 250

	V
	LOD - 100


This range may be extended by sample dilution.
Limit of Detection, Recoveries of Compounds and Uncertainty of measurement:
Agilent 7900 data

	
	
	
	
	Final Effluent (Sutton)
	Trade Discharge
	Untreated Raw
	

	Element
	LOD (µg/l)
	MRV (µg/l)
	
	Low Spike
	High Spike
	Low Spike
	High Spike
	Low Spike
	High Spike
	Dissolved Det

	Al
	9.10
	25
	%recovery
	100.80
	100.51
	103.63
	102.55
	101.96
	103.05
	x

	
	
	
	%rsd
	2.03
	1.63
	2.10
	1.67
	2.67
	1.32
	

	P
	35.3
	36
	%recovery
	103.79
	104.84
	107.35
	106.55
	103.45
	104.72
	x

	
	
	
	%rsd
	2.03
	2.01
	1.44
	1.41
	2.00
	1.72
	

	V
	2.56
	2.6
	%recovery
	101.22
	101.32
	103.64
	102.87
	102.58
	101.85
	-

	
	
	
	%rsd
	2.41
	1.90
	2.77
	1.43
	2.60
	1.86
	

	Cr
	0.227
	0.81
	%recovery
	98.85
	99.08
	93.18
	99.59
	100.07
	100.86
	x

	
	
	
	%rsd
	1.66
	1.88
	1.78
	1.85
	1.53
	1.68
	

	Mn
	0.646
	0.81
	%recovery
	98.94
	99.06
	100.18
	100.54
	99.07
	100.00
	x

	
	
	
	%rsd
	1.59
	1.63
	1.75
	1.31
	1.46
	1.21
	

	Fe
	9.61
	31
	%recovery
	96.73
	98.11
	97.36
	99.72
	98.09
	89.43
	x

	
	
	
	%rsd
	1.66
	1.93
	1.63
	1.57
	1.99
	2.16
	

	Co
	0.114
	0.18
	%recovery
	99.89
	99.57
	101.89
	101.61
	100.86
	100.64
	-

	
	
	
	%rsd
	2.03
	2.11
	1.51
	1.79
	1.57
	1.40
	

	Ni
	0.539
	2.5
	%recovery
	96.44
	96.81
	100.35
	99.94
	96.07
	96.78
	x

	
	
	
	%rsd
	2.08
	2.03
	1.79
	1.67
	1.64
	1.58
	

	Cu
	1.18
	2.5
	%recovery
	98.30
	98.73
	102.22
	100.94
	93.71
	98.64
	x

	
	
	
	%rsd
	1.91
	1.85
	2.46
	1.58
	2.82
	2.32
	

	Zn
	7.46
	7.5
	%recovery
	100.66
	104.64
	99.64
	108.18
	104.23
	97.25
	x

	
	
	
	%rsd
	2.77
	1.92
	2.98
	4.61
	5.56
	1.45.
	

	As
	0.201
	0.80
	%recovery
	100.53
	100.86
	101.60
	100.91
	97.71
	97.34
	x

	
	
	
	%rsd
	1.66
	1.72
	1.41
	1.57
	1.47
	1.05
	

	Se
	0.272
	0.80
	%recovery
	99.45
	99.40
	98.79
	99.65
	-
	-
	x

	
	
	
	%rsd
	2.30
	1.72
	1.48
	1.51
	-
	-
	

	Mo
	0.206
	1.9
	%recovery
	100.49
	100.87
	100.77
	101.08
	96.14
	98.15
	-

	
	
	
	%rsd
	1.85
	1.89
	1.84
	1.68
	3.24
	2.84
	

	Cd
	0.0837
	0.12
	%recovery
	100.92
	100.69
	104.44
	103.09
	102.67
	102.33
	x

	
	
	
	%rsd
	1.68
	1.69
	1.53
	1.30
	1.44
	1.17
	

	Tl
	0.196
	0.20
	%recovery
	99.23
	99.36
	101.28
	101.57
	99.93
	100.17
	-

	
	
	
	%rsd
	1.64
	1.56
	1.33
	1.28
	1.31
	0.90
	

	Pb
	0.109
	0.81
	%recovery
	100.42
	100.69
	103.40
	104.06
	101.55
	102.55
	x

	
	
	
	%rsd
	1.60
	1.59
	3.25
	1.29
	1.16
	0.95
	


References:

In house method based on SCA bluebooks

ISO 17294-2:2016: Water quality — Application of inductively coupled plasma mass spectrometry (ICP-MS) - Part 2: Determination of selected elements including uranium isotopes
Agilent ICP-MS: ISIS 3 (Integrated Sample Introduction System).

Agilent ICP-MS: MassHunter Workstation User Guide.

Agilent ICP-MS: Familiarization Guide

Agilent 7800/7900 ICPMS: Hardware Maintenance Manual

Agilent SPS 4 Autosampler: User’s Guide

ISO 17294-2:2003 Water quality – Application of inductively coupled plasma mass spectrometry (ICP-MS) – Part 2: Determination of 62 elements.
ISO 5667-3:2024 Water quality. Sampling. Preservation and handling of water samples
	WWC46
	Last Updated: 21/05/26
	Page 1 of 1



[image: image1.jpg]