	METHOD STATEMENT
	[image: image1.jpg]N







Determinand:
Determination of Dissolved and Total Metals: The elements covered are aluminium, boron, calcium, cadmium, cobalt chromium, copper, iron, potassium, magnesium, manganese, sodium, nickel, phosphorus, lead, thallium, vanadium, zinc and sulphate.
Matrix:

Sample Type: Final effluents, trade discharges, crude sewage and other similar waste water samples

Principle of Method:

The method describes a technique for the simultaneous multi – element determination of metals in solution.  Samples are acidified and digested at elevated temperatures to ensure any particulate matter is brought into solution.  The acidified samples are then analysed by the ICP-OES (inductively coupled plasma optical emission spectroscopy) instrument to determine the concentration of metals present.  Excitation of the sample within the 6,000°C plasma causes ionisation of atoms, which in turn causes the emission of electromagnetic radiation at specific wavelengths for each element.  The intensity of the emission is measured and quantified by comparison against standards with known concentrations of elements.
Sampling and Sample Preparation:

Samples are normally taken in 1L PET bottles. Samples are stored at 3 ± 2°C until ready for preparation.

Samples should be prepared within 7 days of sampling (Coventry in-house data).

Prepared samples are stable for one month. (ISO 5667-3:2024). Prepared samples are stored at room temperature.
Interferences
Careful choice of plasma conditions, analytical wavelengths and concentration ranges are all designed to minimise interference.
Performance of Method:

Range of Application:

	Element
	Units
	Range Optima
	Range Agilent

	Al 
	mg/l
	LOD – 50
	LOD – 50

	Ca
	mg/l
	LOD – 250
	LOD – 250

	Fe
	mg/l
	LOD – 50
	LOD – 50

	K
	mg/l
	LOD – 100
	LOD – 100

	Mg
	mg/l
	LOD – 150
	LOD – 150

	Mn
	mg/l
	LOD – 50
	LOD – 50

	Na
	mg/l
	LOD – 200
	LOD – 200

	Cd
	mg/l
	LOD – 1
	LOD – 1

	Co
	mg/l
	LOD – 1
	LOD – 1

	Cr
	mg/l
	LOD – 10
	LOD – 10

	Cu
	mg/l
	LOD – 10
	LOD – 10

	Ni
	mg/l
	LOD – 10
	LOD – 10

	Pb
	mg/l
	LOD – 10
	LOD – 10

	V
	mg/l
	LOD – 1
	LOD – 1

	Zn
	mg/l
	LOD – 10
	LOD – 5

	P
	mg/l
	LOD – 10
	LOD – 10

	B
	mg/l
	LOD – 10
	LOD – 10

	Tl
	mg/l
	LOD – 5
	LOD – 5

	SO4
	Mg/l
	LOD - 1000
	LOD - 1000


All analytical ranges may be extended by sample dilution.
Limit of Detection, Recoveries of Compounds and Uncertainty of measurement:
OPTIMA
	
	Units
	LOD 
	Reporting Limit 
	Max. Allowable Blank
	Final Effluent Spike Recovery %
	Spike RSD %
	Uncertainty %

	Aluminium
	mg/l
	0.0809
	0.09
	0.09
	99.08
	3.30
	7.52

	Boron
	mg/l
	0.0408
	0.041
	0.041
	97.82
	3.12
	8.42

	Calcium
	mg/l
	2.7213
	2.8
	2.8
	95.61
	3.69
	11.78

	Cadmium
	mg/l
	0.0048
	0.005
	0.005
	99.33
	2.94
	6.56

	Cobalt
	mg/l
	0.0046
	0.005
	0.005
	101.08
	3.19
	7.45

	Chromium
	mg/l
	0.0223
	0.023
	0.023
	98.77
	3.47
	8.18

	Copper
	mg/l
	0.0256
	0.026
	0.026
	98.14
	3.41
	8.68

	Iron
	mg/l
	0.5626
	0.6
	0.6
	98.14
	3.67
	9.19

	Potassium
	ug/l 
	907.7
	1000
	1000
	95.78
	3.68
	11.58

	Magnesium
	mg/l
	0.5918
	0.6
	0.6
	97.18
	3.32
	9.46

	Manganese
	mg/l
	0.0225
	0.091
	0.091
	97.28
	3.44
	9.60

	Sodium
	ug/l
	1791.5
	1800
	1800
	94.45
	3.90
	13.36

	Nickel
	mg/l
	0.0086
	0.061
	0.061
	97.84
	3.86
	9.89

	Phosphorus
	mg/l
	0.4729
	0.48
	0.48
	89.65
	3.24
	16.83

	Lead
	mg/l
	0.2838
	0.29
	0.29
	98.38
	3.64
	8.90

	Thallium
	µg/l
	98.6
	100
	100
	97.57
	4.03
	10.50

	Vanadium
	mg/l
	0.0134
	0.014
	0.014
	101.69
	3.09
	7.88

	Zinc
	mg/l
	0.0698
	0.07
	0.07
	97.19
	3.31
	9.43

	Sulphate
	mg/l
	17.22
	18
	18
	98.83
	3.11
	12.28


For Urban Waste Water Regulations, an assessment of the method performance with different forms of phosphorus needs to be made. This is summarised below.

Potassium dihydrogen orthophosphate:
Recovery 97.56%

Sodium pyrophosphate:
Recovery 97.38%

Sodium hexametaphosphate:
Recovery 102.43%

B- Glycerophosphate:
Recovery 99.27%
AGILENT
	
	Units
	LOD
	Reporting Limit
	Property
	Crude
	Final
	Trade
	Filtered
	Uncertainty %

	
	
	
	
	
	Low
	High
	Low
	High
	Low
	High
	Low
	High
	

	Al
	mg/l
	0.0344
	0.09
	% Recovery
	99.64
	99.79
	103.77
	100.44
	101.96
	101.60
	100.00
	100.14
	8.006

	
	
	
	
	%RSD
	3.15
	2.75
	2.17
	2.39
	2.57
	2.24
	3.10
	2.75
	

	B
	mg/l
	0.0137
	0.041
	% Recovery
	103.44
	102.34
	104.88
	102.29
	107.41
	109.12
	N/A
	N/A
	11.696

	
	
	
	
	%RSD
	3.10
	2.76
	2.41
	2.00
	2.49
	2.54
	N/A
	N/A
	

	Ca
	mg/l
	2.629
	2.8
	% Recovery
	93.28
	96.72
	100.74
	97.87
	99.61
	101.31
	95.97
	104.59
	21.662

	
	
	
	
	%RSD
	3.02
	3.00
	2.19
	2.44
	2.27
	2.12
	2.68
	2.67
	

	Cd
	mg/l
	0.00196
	0.005
	% Recovery
	106.07
	101.72
	105.76
	102.55
	106.74
	102.58
	106.38
	102.77
	11.67

	
	
	
	
	%RSD
	2.99
	2.63
	2.10
	1.74
	2.41
	2.06
	2.60
	2.56
	

	Co
	mg/l
	0.00320
	0.005
	% Recovery
	104.67
	102.54
	104.12
	101.87
	103.43
	101.71
	N/A
	N/A
	8.985

	
	
	
	
	%RSD
	3.72
	2.46
	2.76
	2.36
	2.58
	2.19
	N/A
	N/A
	

	Cr
	mg/l
	0.00702
	0.023
	% Recovery
	101.21
	102.28
	101.59
	102.07
	118.15
	102.98
	101.89
	102.26
	13.728

	
	
	
	
	%RSD
	2.88
	2.81
	1.96
	2.00
	2.17
	2.29
	2.63
	2.64
	

	Cu
	mg/l
	0.0100
	0.026
	% Recovery
	103.67
	102.36
	100.98
	102.65
	83.31
	98.93
	101.26
	102.81
	13.096

	
	
	
	
	%RSD
	2.64
	2.39
	2.25
	1.82
	2.13
	2.01
	2.78
	2.52
	

	Fe
	mg/l
	0.0395
	0.6
	% Recovery
	103.13
	100.91
	104.01
	99.98
	100.22
	101.22
	98.53
	99.29
	7.7

	
	
	
	
	%RSD
	2.90
	2.81
	2.24
	2.64
	2.73
	2.17
	3.30
	2.71
	

	K
	µg/l
	1037
	1040
	% Recovery
	100.96
	98.72
	97.60
	98.67
	86.63
	97.38
	100.79
	99.86
	11.181

	
	
	
	
	%RSD
	2.67
	1.75
	2.02
	1.65
	1.98
	1.88
	2.72
	2.72
	

	Mg
	mg/l
	0.400
	0.6
	% Recovery
	105.31
	100.62
	106.36
	101.14
	103.77
	102.53
	103.82
	101.13
	9.455

	
	
	
	
	%RSD
	2.98
	2.42
	2.27
	2.36
	2.35
	1.84
	2.69
	2.54
	

	Mn
	mg/l
	0.0275
	0.091
	% Recovery
	102.85
	100.72
	100.59
	100.03
	99.44
	101.24
	99.93
	99.94
	6.93

	
	
	
	
	%RSD
	2.98
	2.88
	2.13
	2.67
	2.32
	2.33
	2.71
	2.72
	

	Na
	µg/l
	3096
	3100
	% Recovery
	100.94
	96.41
	99.75
	97.33
	100.23
	94.64
	100.51
	101.38
	8.097

	
	
	
	
	%RSD
	3.31
	2.38
	2.60
	1.31
	2.47
	2.70
	2.86
	2.74
	

	NI
	mg/l
	0.009
	0.061
	% Recovery
	102.03
	100.04
	99.94
	100.05
	98.53
	100.27
	101.41
	101.5
	7.276

	
	
	
	
	%RSD
	2.81
	2.09
	2.47
	2.01
	1.98
	1.97
	2.85
	2.61
	

	P
	mg/l
	0.0328
	0.48
	% Recovery
	105.22
	103.33
	105.51
	102.08
	103.52
	101.69
	107.69
	102.34
	12.304

	
	
	
	
	%RSD
	2.45
	1.43
	2.48
	1.83
	2.03
	2.12
	2.32
	1.64
	

	Pb
	mg/l
	0.0210
	0.29
	% Recovery
	103.24
	102.81
	101.37
	101.73
	99.9
	100.35
	102.47
	102.09
	6.868

	
	
	
	
	%RSD
	2.00
	1.63
	2.51
	1.95
	2.19
	2.24
	2.28
	1.61
	

	Tl
	µg/l
	14.89
	100
	% Recovery
	101.74
	101.31
	100.12
	100.91
	99.10
	102.00
	100.81
	101.53
	6.907

	
	
	
	
	%RSD
	3.12
	2.46
	1.76
	1.58
	2.30
	1.77
	2.91
	2.69
	

	V
	mg/l
	0.00234
	0.014
	% Recovery
	109.09
	103.01
	108.99
	103.15
	108.81
	103.71
	109.05
	104.08
	14.785

	
	
	
	
	%RSD
	2.92
	2.84
	2.10
	1.81
	2.02
	1.94
	2.49
	2.38
	

	Zn
	mg/l
	0.0418
	0.07
	% Recovery
	107.58
	109.11
	108.93
	107.91
	108.7
	108.86
	107.87
	108.11
	17.638

	
	
	
	
	%RSD
	1.65
	2.52
	2.63
	2.14
	1.82
	3.19
	2.04
	1.93
	

	SO4
	mg/l
	8.336
	18
	% Recovery
	105.77
	103.34
	105.88
	102.21
	103.7
	100.87
	106.1
	102.58
	10.675

	
	
	
	
	%RSD
	1.82
	1.43
	2.49
	1.72
	2.04
	1.99
	2.56
	1.48
	


For Urban Waste Water Regulations, an assessment of the method performance with different forms of phosphorus needs to be made; recoveries are summarised below:

Potassium dihydrogen orthophosphate:
Recovery: 107.57%

Sodium pyrophosphate:
Recovery: 104.19%

Sodium hexametaphosphate:
Recovery: 101.20%

B- Glycerophosphate:
Recovery: 103.39%

References:

In house method based on SCA bluebooks, and Perkin Elmer Optima 7100, 7200 and 7300 series Hardware Guide manual and Agilent 5900 user guide.
ISO 5667-3:2024 Water quality. Sampling. Preservation and handling of water samples
	WWC08
	Last Updated: 24/06/26
	Page 2 of 2



[image: image1.jpg]